Упутство за ауторе радова 

HIDROGELI NA BAZI POLIAKRILNE KISELINE U SVRHU SMANJENJA BIOADSORPCIJE NA SILIKONSKIM KONTAKTNIM LEĆAMA
Z. Ademovic1,*, S. Maric2, P. Kingshott3, Z. Ilickovic1
1Univerzitet u Tuzli, Tehnološki Fakultet, 
Univerzitetska 8, 75000 Tuzla, e-mail: zahida.ademovic@untz.ba 
2Univerzitet u Tuzli, Prirodno-Matematički Fakultet, 
Univerzitetska 6, 75000 Tuzla, e-mail: snjezana.maric@untz.ba 
3Swinburne University of Technology, Inženjerstva i Industrijskih nauka , Hawthorn, 

Australija, e-mail: pkingshott@swin.edu.au

HYDROGELS FROM POLYACRYLIC ACID FOR REDUCTION OF BIOADHESION ON SILICONE CONTACT LENSES
Z. Ademovic1,*, S. Maric2, P. Kingshott3, Z. Ilickovic1
1University of Tuzla, Faculty of Techonolgy, 
Univerzitetska 8, 75000 Tuzla, e-mail: zahida.ademovic@untz.ba 
2University of Tuzla, Faculty of Science, 
Univerzitetska 6, 75000 Tuzla, e-mail: snjezana.maric@untz.ba 
3Tehnološki Univerzitet Swinburn, Hawthorn, 

Australia, e-mail: pkingshott@swin.edu.au

Sažetak: Kontaktne leće imaju dva ograničenja: niska propusnost kisika i taloženje proteina i lipida. Kako bi se spriječila adhesija bioloških molekula na površinu materijala, materijal mora biti potpuno inertan na sve biološke reakcije. Ovo je moguće postići modificiranjem površine materijala. Kako bi se poboljšala udobnost nošenja kontaktnih leća, površinu mora biti vrlo hidrofilna i lubriosna. U ovom radu površina silikonskih kontaktnih leća modificirana je putem plazma inducirane kopolimerizacije akrilne kiseline. Na taj način stvorena je veoma hidrofilna površina, a pored toga karboksilne skupine koje su stvorene na površini predstavljeju idealnu platformu za naknadno modificiranje sa polietilenglikolom. Svaki korak modifikacije analiziran je putem XPS i mjerenja kontaktni ugla. Adsorpcija proteina lisozima na površinu silikonskih leća analizirana je sa surface-MALDI-TOF-MS i XPS. Nakon inkubacije sa lisozimom, analiza je pokazala smanjenje adsorbiranog lisozima na hidrogel modificiranim kontaktnim lećama. Površinska modifikacija silikonskih kontaktnih leća sa PEG je prihvatljiva metoda za smanjenje adsorpcije proteina  na kontaktne leće.
Ključne riječi: hidrogeli, PEG, adsorpcija proteina, kontaktne leće
Abstract: Contact lenses suffers from two limitations: low oxygen permeability and deposition of protein and lipids. In order to prevent bioadhesion surface must be completely inert to all biological reactions and to achieve this surface properties must be tailored. Also, to improve comfort surface must be highly wettable and lubrious. In this work surface of silicon contact lenses was modified by plasma induced copolymerization of acrylic acid. A wettable surface was generated and in addition carboxyl groups that were created on the surface provided the ideal reactive platform for subsequent grafting of polyethylene glycol. Each surface modification step was analysed by XPS and contact angle measurements. Lysozyme adsorption on modified silicon contact lenses was analysed by surface-MALDI-ToF-MS and XPS. After incubation with lysozyme, surface-MALDI-TOF-MS and XPS analysis showed a reduction of adsorbed lysozyme on hydrogel modified contact lenses. Surface modification of silicon with PEG is a method for reduction of protein adsorption on contact lenses.
Key words: hydrogels, PEG, protein adsorption, contact lenses 

1. INTRODUCTION

Formation of protein deposits on the lens from the tear fluid is one of the problems encountered with contact lense use. The deposit cloud the lens, cause the wearer discomfort, reduce oxygen permeability and may cause inflammatory conditions. The development of lens materials taht exibits minimal protein adsorption is therefore desirable. The rate and amount of adsorbed protein are dependent on physico-chemical properties of the polymeric surface such as wettability and surface charge density. Therefore surface modification procedures are developed to create ultrathin protein-repellent nanostructured and biologically functionalized interfaces at the implant surfaces. Silicone is a material that is widely used in many applications, but as material for contact lenses it has a limited application. The reason is high hydrophobic nature of silicone which leads to inadequate material properties and performance of contact lenses when in contact with tear film. All hydrophobic surfaces when in contact with biological fluids tend to adsorb and deposit lipids and proteins on the surfaces, which influence biocompatibility of whole material [1]. For contact lenses it means that wearing comfort is affected, it can lead to dryness of natural lens and it can reduce oxygen permeability. For the optimization of the biocompatibility of silicone contact lenses a specific surface design is needed in which aims at the prevention of nonspecific protein adsorption and bacterial adhesion. The specific surface modification of silicone however requires the introduction of suitable functional groups. Surface modification is achieved without influencing bulk properties of the material.
Grafting of poly(ethylene glycol) (PEG) is a common strategy for reducing the non-specific adsorption of biomolecules on surfaces [2-3]. Several theories – both qualitative and quantitative – have been proposed to explain the effectiveness of PEG as a component of protein-resistant surfaces [4]. An essential finding of these theories is that the PEG layer might provide an interfacial barrier to prevent the protein from interacting with the underlying substrate. The protein rejecting capability of PEG coatings is expected to depend on a range of parameters, including PEG molecular weight, interfacial PEG chain density, polymer chain architecture, etc, the interfacial graft density proving to be the most crucial property of a PEG layer for minimizing protein adsorption [5-6]. 

In this work the functionalization of the silicone surface is achieved by plasma-induced graft-co-polymerisation of acrylic acid (AAc). Plasma-induced grafting of hydrogels like polyacrylic acid offers a possibility to obtain permanent minimization of the surface energy of hydrophobic polymer surfaces [7]. Acrylic acid was chosen as the hydrogel building monomer to generate a stable, hydrophilic surface with the additional possibility of using the carboxylic groups for further modifications. In this work a hydrogel layer of polyethylenimine (PEI) is attached to the PAAc modified silicone surface. Additionally, metoxy-aldehyde polyethylene glycol (M-PEG-ald) is grafted on the surface to prevent nonspecific adsorption of proteins.
All modification steps were verified by means of X-ray photoelectron spectroscopy (XPS) and contact angle measurements. Protein adsorption studies were carried out by means of surface MALDI-ToF-MS. Modified PEG graft-modified silicone surface can be used as material with a nonfouling surface while still bearing the excellent bulk properties of silicone. 
2. MATERIALS AND METHODS
2.1. Materials
Referent material was Sylgard® 184, a two component silicone elastomeric kit obtained from Dow Corning. Acrylic acid was obtained from Fluka (Germany) and distilled prior use. Methoxy-terminated formyl-PEG (M-PEG, molecular weight 5000) were purchased from Shearwater Polymers. Lysozime and tear artificial fluid were purchased from Sigma.
2.2. Chemical functionalisation
Graftcopolymerisation: Plasma treatment was carried out with a microwave plasma unit of Eltron (Germany). PVDF films were argon plasma treated at a plasma power of 900 W for 30 s. Subsequently, the air-exposed PVDF-films were immersed in an aqueous solution containing 20% (v/v) of AAc monomer and heated at 90 °C to initiate graft polymerisation. The PAAc-grafted films were rinsed with distilled water for 24 h to remove nongrafted monomers and PAAc-homopolymers and stored in distilled water until use. 
The carboxyl end groups of the grafted silicone were activated with 0.1 M EDC in sodium phosphate buffer pH 4.8 for 20 min at room temperature. Afterwards coupling of polyethylenimine was carried out in carbonate buffer pH 9.4 for 2 h at room temperature.

M-PEG was grafted onto aminated surface by reductive amination, using NaCNBH3 as the reducing agent for the intermediate Schiff’s base. The grafting was performed under marginal solvation conditions (´cloud point`). 1 mg/ml M-PEG was dissolved in 0.1 M sodium phosphate buffer at pH 6.3 containing 11% (w/v) K2SO4. The reaction was carried out at 60°C [8]. 
All surfaces were incubated in 1 mg/ml lysozyme in PBS for 1 hour at 37°C. The samples were rinsed 3x with water and analysed immediately with MALDI-ToF-MS.
2.3. Physical and chemical surface characterization
All X-ray photoelectron spectra (XPS) were recorded on an X-Probe( 206 spectrometer (Surface Science Instruments, Mountain View, CA). X-ray source is an aluminium anode whic produces AlK( X-rays at 1486.6 eV. The binding energies are referenced to hydrocarbon at 285.0 eV. The emission angle of electrons was set at 55° with respect to the sample normal, which results in an information depth of about 6 nm. 

Contact angles were measured using the captive bubble method with pure water at room temperature on a G40 system (Krüss, Hamburg, Germany).

Surface-MALDI-ToF-mass spectra were obtained using a BRUKER BIFLEXTM III MALDI time-of-flight mass spectrometer (Bruker-Franzen Analytik GmbH, Bremen, Germany) equipped with a nitrogen laser. In this experiments a small sample piece cut from silicone and modified silicone were placed onto the stainless steel MALDI sample holder. Sinapinic acid in a 0.1% solution of trifluoroacetic acid in acetonitrile/water was applied onto the sample surface and the solvent was left to evaporate before the sample holder was inserted into the spectrometer.

3. RESULTS AND DISCUSSION
Plasma technology is applied for activation and functionalisation of silicone surface. Figure 1 shows a schematic illustration of the silicone surface modifications by graft-co-polymerisation of AAc. Upon argon plasma treatment radicals are formed on the polymer surface, which react with oxygen in contact with air and form hydroperoxides. The subsequent thermal decomposition of the hydroperoxides produces secondary radicals that are able to initiate graft-co-polymerisation of acrylic acid (AAc) to produce carboxylic groups on the silicone surface (silicone-PAAc). In a following step polyethyleneimine (PEI) was grafted on the carboxyl groups on the silicone surface to create amino-groups (silicone-NH2). Subsequently, M-PEG-ald was grafted onto the aminated surface. The graft density and length of PEG chains are essential for protein repellent properties. The high density of PEG chains is achieved by reaction of methoxy-polyethylenglykol-aldehyde (mPEG-ald) with amino groups of PEI layer at ´cloud point´ reaction conditions (60°C, 11% K2SO4). Resulting surface (silicone-PEG) has protein repellent properties.
[image: image1.png]OOH radical 3 — COOH

polymerisation
; . argon plasma OOH A COOH
. 2l exposwe -
OOH ° — COOH
OOH \)LOH — COOH
silicone Silicone with generated Arc Silicone-PAAC

hydroperoxide groups




[image: image2.png]= sy sy
8 8 8
© O O

NH2 PEI-NH2 ! ! l
EDC/NHS

Silicone-PAAC Silicone-NH,




[image: image3.png]@

o %

¢144
2y

prote

N

mPEG-ald

Oxto A2

e-PEG

Silicon




Figure 1: Activation and functionalisation of silicone surface

The hydrophilicity of the surfaces, reflected by the contact angle is expected to be significantly reduced for the modified surfaces. The results of contact angle measurements by the captive bubble method confirmed these expectations. Unmodified silicone has a contact angle of 83±5°, which indicates a high hydrophobe surface. After graft-co-polymerisation of acrylic acid, an increase in wettability with water is achieved, contact angle was considerably reduced to 29±2°. A high hydrophilic surface is produced due to the carboxylic groups. After grafting of PEG chains contact angle is further reduced to 19±4°.

Quantitative information about the chemical composition of the outermost (10 nm) surface layer of the unmodified and modified silicone surfaces has been obtained by means of XPS. All modifications steps were reflected in clear alterations of the surface element composition. XPS data showing the elemental composition as well as the different carbon species are listed in Table 1.
High resolved C1s-spectra from surfaces showed three different carbon species originating from aliphatic, ether and carboxyl carbon. The results showed the successful grafting of acrylic acid on the plasma activated silicone surface. After graft copolymerisation of PAAc, newly formed O(C=O groups are detected at 289.1 eV, whereas silicium content decreased from 27.5 to 21.6 atom-%. Covalent grafting of a PEI layer results in an introduction of amino groups on the silicone-PAAc surface. As shown in Table 1 surface nitrogen content of 2.3 atom-% was registered and it is an indicator for PEI present on the surface. Also, the carbon position at 286.5 eV is attributed to C-N carbon of amino groups overlaps with C-O signals and increased from 2.0 to 3.5 atom-%. 

The efficiency of grafting the M-PEG chains onto the silicone-PEI surface was derived from the ratio of the XPS intensities of the ether carbon atom (originating from PEG) and alkyl carbon atom (originating from silicone-PEI). Deconvolution of the C1s-XPS of the silicone-PEG surface (Table 1) shows two dominant contributions with respective binding energies of 285.0 eV and 286.5 eV. The strong 286.5 eV contribution from the ether carbon atom of M-PEG-ald indicates a successful binding reaction between M-PEG-ald and the PEI coated surface. Pure PEG samples exhibit a single C1s peak centered at 286.5 eV.
The XPS data of Table 1 also show that M-PEG-ald coatings invariably is thinner than the information depth of the XPS measurements as the analyses does not match the theoretical composition of pure PEG (a single C-O peak and no N from underlaying amine PEI layer). Thus, XPS represent superpositions from the M-PEG-ald coatings and the underlying layers.
Table 1: Untreated and modified silicone surfaces characterized by means of XPS for elemental composition and binding energy
	Surface
	Si 

atom-%
	Carbon (C1s)
atom-%
	O

atom-%
	N

atom-%

	
	
	285.0 eV
	286.5 eV
	289.1 eV
	
	

	
	
	C-C

C-H, C-Si
	C-O

C-N
	O-C=O

N-C=O
	
	

	Silicone
	27.5
	45.5
	1.9
	
	25.2
	

	Silicone-PAAc
	21.6
	41.3
	2.0
	5.4
	28.9
	

	Silicone-PEI
	18.4
	51.5
	3.5
	1.1
	23.9
	2.3

	Silicone-PEG
	13.6
	29.8
	27.9
	1.2
	27.1
	0.4


After all modification steps a relatively high content of silicium is still registered. Reasons for detected silicium are reorganization effects on the surface. In order to minimize free surface energy the hydrophobic methyl groups settle outwards, as the XPS spectra are measured from dried samples in high vacuum.
The ability of M-PEG layer to repel in vitro lysozyme adsorption was assessed by surface-MALDI-ToF mass spectrometry. Surface-MALDI-ToF-MS detects small amounts of adsorbed material [9] but is difficult to quantify. Lysozyme is a relatively small protein with isoelectric point at 11.35 and with a net positive charge at pH 7.4 (physiological condition). It is one of mayor proteins in human tear fluids. Thus, its adsorption on a hydrogel material should be promoted as the density of negatively charged groups increases. On the other hand, lysozyme uptake into positively charged hydrogel matrices should be reduced due to charge repulsion effects. 
Due to the high hydrophilic nature of silicone, unmodified silicone surface strongly adsorbs lysozyme (Figure 2). The peaks observed at m/z values of 14229 and 7072 are assigned to the protonated molecular ion (M+H)+ and the doubly charged molecular ion (M+2H)2+ of lysozyme. Carboxylic groups from silicone-PAAc surface are negatively charged at pH 7.4. Electrostatic attraction between the negatively charged carboxy groups of ionised polyacrylic acid and the overall positive charge of lysozyme is the cause of lysozyme uptake by such a hydrogel [10]. Figure 2 shows that PAAc layer on a silicone film enhanced the adsorption of lysozyme whereas grafting of the cationic PEI yields a surface that strongly rejects lysozyme.
Surface–MALDI-ToF mass spectrum measured on the silicone-PEG coating after lysozyme adsorption shows no peaks of adsorbed lysozime. It indicates that grafting density of the PEG chains is sufficient to prevent lysozime adsorption.
	[image: image4.png]1000 |

1 —— Silkon
2 —— Silkon-PAAC
3 ——Silkon-PAACPEI

T
10000

T
15000
mlr




	[image: image5.png]100D —— silikon-PEG

° T T T 1
10DoD 15000 20000 25000
——







Figure 2: Surface-MALDI-TOF-MS spectra of lysozyme adsorbed on the silicone (1), silicone-PAAc (2) and silicone-PEI (3) surfaces (left) and on the silicone-PEG surface (right)

XPS results of adsorption of 1 mg/ml lysozyme (PBS-buffer pH 7.4 for 1 hours at 37°C) on the nonmodified and modified silicone surfaces are shown at Figure 3. 
[image: image6.png]Atom %

50

40

30

51,8 51,5
474
oc
27 p9,3 [=]e)
75,2
P78 1233 218 an
194 180 |osi
0.0 21 1,1 1,2
Silicone Silicone+Lys Silicone-PEG Silicone-PEG+Lys




Figure 3: XPS results of 1 mg/ml lysozyme adsorption on silicone and silicone-PEG surfaces
Nitrogen content of 2,1 atom-% after adsorption of lysozyme is measured on nonmodified silicone surface. As silicone does not contain nitrogen, this can originate only from protein lysozyme, which indicated a strong adsorption of lysozyme onto silicone due to hydrophobic nature of the surface 

After incubation of silicone-PEG surface with lysozyme, there is no increase of nitrogen content detected. PEG modified silicone surfaces inhibited protein adsorption onto surface.

4. CONCLUSIONS

In the present study plasma graft-co-polymerisation was used to produce hydrogel coatings on the silicon contact lenses that can minimise protein adsorption. In the first step carboxylic groups were introduce to a silicone substrate by graft-co-polymerisation of AAc. After attachment of PEI to the carboxylic groups, linear mPEG chains were grafted on the surface. Maximal grafting density of mPEG chains onto surface was achieved by grafting at the lower critical solution temperature of PEG. All surfaces modification steps were analyzed using contact angle measurements, and XPS. Protein adsorption measurements were performed with lysozyme using surface MALDI-TOF MS and XPS. Results showed that the high density of PEG chains on the silicone contact lenses surfaces can be produced which inhibits protein adsorption.
5. REFERENCES
[1] B.D. Ratner, A.S Hoffman, F.J Shoen, J.E Lemons, editors. Biomaterials science: Introduction to Materials in Medicine, 2nd edition. Elsevier Academic Press, London, UK 2004. 
[2] J.M. Harris, Poly(Ethylene Glycol) Chemistry: Biotechnical and Biomedical Applications, Plenum Press: New York, 1992
[3] S.J. Sofia, V. Premnath, E.W. Merrill, Poly(ethylene oxide) Grafted to Silicon Surfaces: Grafting Density and Protein Adsorption, Macromol. 31 (1998) 5059-507

[4] S.I. Jeon, J.H. Lee, J.D. Andrade, P.G. de Gennes, Protein-surface interactions in the presence of polyethylene oxide: I. Simplified theory, J. Colloid Interface Sci. 142 (1991) 149-158
[5] I. Szleifer, Protein Adsorption on Surfaces with Grafted Polymers: A Theoretical Approach, Biophys. J. 72 (1997) 595-612
[6] J. Groll, Z. Ademovic, T. Ameringer, D. Klee, M. Moeller, Comparison of coatings from reactive star shaped PEG-stat-PPG prepolymers and grafted linear PEG for biological and medical applications, Biomacromolecules, 6 (2005) 956-962
[7] D. Klee, Z. Ademovic, H. Hoecker, A.K. Bosserhoff, G. Maziolisc, H.J. Erli, Surface modification of polyvinylidene fluoride to improve the cell adhesion, Biomaterials 24 (2003) 3663-3670.

[8] Z. Ademovic, D. Klee, P. Kingshott, H. Hoecker, Minimization of protein adsorption on polyvinylidene fluorid, Biomolecular Engineering, 19 (2002) 177-182
[9] P. Kingshott, A.W. Heather, S. John, H. Griesser, Direct Detection of Proteins Adsorbed on Synthetic Materials by Matrix-Assisted Laser Desorption Ionization–Mass Spectrometry, Anal. Biochem. 273 (1999) 156-162
[10] R.A. Sack, B. Jones, A. Antignani, R. Libow, H. Harvey, Specificity and biological activity of the protein deposited on the hydrogel surface. Relationship of polymer structure to biofilm formation, Invest. Ophthalmol. Vis. Sci. 28 (1987) 842-849






