DETERMINATION OF TOTAL CHLORINE IN WATER BODIES BY FIBER OPTIC SENSOR

Obrovski Boris1*, Bajić Jovan2, Mihajlović Ivana1, Vojinović Miloradov Mirjana1, Batinić Branislav2, Živanov Miloš2

1University of Novi Sad, Faculty Of Technical Sciences, Department Of Environmental Engineering, Trg Dositeja Obradovića 6, 21000 Novi Sad, Serbia (e mail: borisobrovski@uns.ac.rs)

2University of Novi Sad, Faculty Of Technical Sciences, Department of Power, Electronic and Telecommunication Engineering, Trg Dositeja Obradovića 6, 21000 Novi Sad, Serbia

Water quality monitoring program needs to be improved by new methods for the detection of chemical species in the examined water body. Fiber optic sensor (FOS) represents innovative method for monitoring of total chlorine in surface water and water from the swimming pool. FOS absorbs light and based on the color of the sample determines concentration of analytes of interest. FOS converts RGB (Red-Green-Blue) color model in HSV (Hue-Saturation-Value) color model and based on S value determines concentration of total chlorine in water media. FOS was applied for examination concentration level of Cl2 in samples of river Danube surface water and the water from the swimming pool. Results obtained by standard laboratory analyses (UV-Vis spectrophotometer) are compared with results obtained by FOS in order to demonstrate reproducibility and efficiency of the FOS device. Implemented sensor confirmed to be suitable in laboratory analyses under controlled conditions as substitute for expensive analytical equipment.

Key words: Fiber optic sensor, total chlorine, surface water, swimming pool

INTRODUCTION

Wastewater discharged from anthropogenic sources and natural processes disturb the quality of water bodies. Ecological and chemical status of environment, especially aquatic system which is used for human activities require high-quality monitoring program therefore of high negative impact on humans and aquatic life. Water from the river Danube can be used for irrigation, industry, human activities, and the most important application is for production of drinking water. Public swimming pools are used for recreation and it is very important to monitor the quality of pools water to prevent a disturbance of human health.

Standard analytical methods are generally accepted and reliable. However, they demonstrated certain limiting factors and intensive effort on the development of new methods for monitoring of different water bodies has been done. Limiting factors are related to the costly and specific chemicals, the complexity of the analysis, the loss of the analyte in the process of sampling, transportation, extraction and storage and impossibility of obtaining real-time results. There is a possibility of changing the concentrations of contaminant in the examined water body if the sampling is carried out in far locations from the laboratory.

FOS is an innovative method that provides a reliable and real-time data on the current chemical state of the examined water bodies. Advantages of the FOS are: simple to use, low-cost device, small dimensions (which enables measurements where other devices do not have access), resistant to electromagnetic influences and corrosion, possibility to use in high aggressive chemical media, enabling measurements in inaccessible and remote areas,  without electric power at sampling points, etc. Because of all these characteristics, FOS is suitable instrument that can replace robust conventional techniques and can be used for determination where satisfactory equipment does not exist. A large number of FOS have demonstrated good efficiency in laboratory tests at controlled conditions, but the small number of devices is commercialized [1].

Today, a great number of FOS are developed for the measurement of various chemical parameters in selected aqueous media. FOS has been developed to detect very low concentrations of nitrate, which allows FOS usage for measuring the concentration of nitrite in drinking water due to of its high sensitivity [2,3]. FOS which is applied in laboratory controlled conditions is constructed to detect ammonia, nitrates and chromium in the prepared test solutions. Multiparameter device based on FOS has high sensitivity and low limit of detection which provides the ability for different water bodies where the concentrations of observed parameters are extremely low and it is very difficult to determine the concentration by other methods [4]. Multiparameter FOS were constructed for measuring of temperature, pH and potassium ions in drinking water and aqueous solution [5], for measuring of cyanide, phosphate, sulfate, nitrate and nitrite [6] in water samples, for measurement of heavy metal cations Co2+, Cu2+, Ni2+, Fe3+, Cd2+, Zn2+, Pb2+ and Hg2+ in the aquatic environment [7], as well as prototype field device for measuring organic pollutants in groundwater [8]. Multiparameter FOS was calibrated for five parameters (orthophosphate, nitrite, sulfate, total chlorine, and hexavalent chromium) in surface water [9].

This paper presents a color fiber optic sensor used to measure total chlorine in samples of Danube river surface water and in swimming pools. The use of chlorine compounds is widespread as a disinfectant for drinking water and swimming pools. During disinfection processes carcinogenic and toxic organic chlorine by-products could be generated. Concentration of chlorine is changing rapidly. Therefore, it is extremely important to receive real and prompt data of the concentration of the chlorine in examined water bodies. The results obtained by standard analytical methods (UV-Vis spectrophotometer) were compared by the results obtained with the FOS to demonstrate the effectiveness and reproducibility of the device. Investigations were conducted to demonstrate possibility of using the FOS in laboratory controlled conditions as a substitute for common analytical equipment.

MATERIALS AND METHODS

Samples of surface river Danube water for laboratory analysis were collected on the locality of left bank in Danube in Novi Sad (19° 51' 22,987" E, 45° 14' 46,514" N).  Samples of pool water were collected from closed swimming pool in Novi Sad, Serbia. Samples are poured into 1 L glass bottles, stored in hand refrigerator at 4 0C, and transported to the laboratory. Concentrations of total chlorine in water samples were analyzed according to the HACH Method (HACH Method 8167) and measured with UV-VIS spectrophotometer (DR 5000, HACH, Germany). Analyses were carried out in accredited Laboratory for monitoring of landfills, wastewater and air, Department of Environmental Engineering and Occupational Safety and Health in Novi Sad.

Operating principle of implemented FOS is the absorption of light. When the light passes through liquid, certain wavelengths will be transmitted while others are absorbed depending on the color of the tested liquid. Fiber optic sensor detects the color and converts RGB color model in HSV color model. The used sensor determines V and S value and calculates concentration of the parameters of interest, and H value which is used for calculation of wavelength. The sensor consists of three plastic optical fibres (POFs) that emit red, green and blue components mounted around a central optical fibre collecting light reflected from the mirror (Figure 2). Three light-emitting diodes red, green and blue are set to different frequencies. In this way detection of the reflected signal is achieved with only one photodetector and three bandpass filters.

Fig. 1. Applied sensor (FOS)

RESULTS AND DISCUSSION

Total chlorine was determined in surface river Danube water and water from swimming pool. To examine and confirmed the efficacy and applicability of the applied sensor, the results obtained by FOS were compared with results obtained by the standard analytical laboratory methods.

Calibration curves for the sensor were constructed with four different reference samples for total chlorine. Reference sample with the lowest concentration of total chlorine have bright pink color and with increasing concentrations of total chlorine, the color of the sample becomes more intense. Total chlorine concentrations were calculated and determined on the basis of the parameter S, based on calibration curves obtained with the standard solutions. The V values for those parameters were constant and it is not possible to determine their concentrations on the basis of V values with the applied sensor. In Table 1 and Table 2 are presented relative differences for total chlorine between standard laboratory method (UV-Vis spectrophotometer) and sensor.

Table 1. Relative differences between concentrations obtained by UV-Vis spectrophotometer and by FOS in samples of  surface water.

| Measurement | UV-Vis [mg/l] | FOS [mg/l] | Relative difference [%] |
| --- | --- | --- | --- |
| 1 | 0,484 | 0,514 | 6,20 |
| 2 | 0,165 | 0,171 | 3,62 |

Table 2. Relative differences between concentrations obtained by UV-Vis spectrophotometer and by FOS of swimming pool water.

| Measurement | UV-Vis [mg/l] | FOS [mg/l] | Relative difference [%] |
| --- | --- | --- | --- |
| 1 | 0,49 | 0,488 | 0,41 |
| 2 | 0,51 | 0,512 | 0,39 |

Relative differences for observed parameter measured by FOS are lower than 10 %, which confirms the successful usage of optical sensor as new laboratory equipment for monitoring of surface water and water from swimming pool.

Fig. 3. Concentration for total chlorine detected by FOS for surface water based of S value

H value is almost constant for all measurement parameters. Table 3 presents the wavelength at which FOS measures the parameters of interest and the wavelengths at which UV-Vis spectrophotometer measure defined parameters. The wavelengths were calculated based on the equation (1) for implemented sensor:

λ nm=-1.1*H[0]+700								(1)

where λ is wavelength, and H is Hue value.

Table 3. Comparison of wavelengths at which wavelength UV-Vis and FOS determine concentration of chlorine.

| Total chlorine | UV-Vis wavelength [nm] | Calculated corresponding wavelength [nm] | Relative difference [%] |
| --- | --- | --- | --- |
| Wavelength for surface water | 530 | 533 | 0,5 |
| Wavelength for water from swimming pool | 530 | 527 | 0,56 |

After the comparison, it could be concluded that the FOS is reliable and applicable for laboratory measuring of examined parameters in the surface water and water from swimming pool. The FOS presents a „low cost“ and rapid optical device and there is a possibility and it is accetable to replace the expensive standard laboratory equipment.

CONCLUSION

Developed FOS used in this research is capable to measure five parameters in surface water and two parameters in pool water. FOS represents multiparameter device that is capable of measuring six parameters (orthophosphate, nitrite, sulfate, total chlorine, residual chlorine and Cr+6) in surface water and water from swimming pool. Further research should be focused on improving the characteristics of the device to increase the accuracy of the obtained results. The use of quality components for construction of devices such as better selection of quality LED would enable more reliable results for measuring of selected parameters. Low price of components for construction of device enables their use in experimental conditions as a “low-cost” device that can replace the standard expensive equipment used in the laboratory. Developing FOS requires an increase in sensitivity of the device, which would allow the measurement of low concentrations of parameters of interest. Further development of the device includes expanding range of examined parameters observed in the selected aquatic bodies. The future plan is to permit application of FOS for measuring of waste water (color water), landfill leachate and samples of groundwater with sub low concentration of contaminants. The ultimate goal of the research is the construction of field device that will be used in real environmental conditions. Field FOS would be used for in-situ measurements and obtaining reliable and real-time data on the quality of the examined water bodies.

ACKNOWLEDGMENT:

The authors acknowledge for the funding provided by the Ministry of Education, Science and Technological Development of Republic of Serbia under project ‘Development of methods, sensors and systems for monitoring quality of water, air and soil’, number III43008.

LITERATURE

[1] Udd, E.: An overview of fiber-optic sensors, Rev. Sci. Instrum., 66, 1995, 4015-4030.

[2] P.S. Kumar, C.P.G. Vallabhan, V.P.N: Nampoori, V.N.S. Pillai, P. Radhakrishnan: A fibre optic evanescent wave sensor used for detection of trace nitrites in water, Journal of Optics: Pure and applied optics, 4, 2002, 247-250.

[3] N.A. Aljaber, B.R. Mhdi, S.K. Ahmmad, J.F. Hamode, M.M. Azzawi, A.H. Kalad, S.M. Ali: Design and Construction Fiber sensor detection system for water nitrite pollution, Journal of Engineering, 04, 2014, 37-43.

[4] P.S. Kumar: Design and development of fiber optic sensors for trace detection of certain environmental pollutants, Ph D Thesis, 2003.

[5] A. Dybko, W. Wroblewski, E. Rozniecka, K. Pozniakb, J. Maciejewski, R. Romaniuk, Z. Brzozka: Assessment of water quality based on multiparameter fiber optic probe, Sensors and Actuators B, 51, 1998, 208-213.

[6] S. M. Klainer, J. R. Thomas, J. C. Francis: Fiber-optic chemical sensors offer a realistic solution to environmental monitoring needs, Sensors and Actuators B, 11, 1993, 81-86.

[7] N. Malcik, O. Oktar, M. E. Ozser, P. Caglar, L. Bushby, A. Vaughan, B. Kuswandi, R. Narayanaswamy: Immobilised reagents for optical heavy metal ions sensing, Sensors and Actuators B, 53, 1998, 211-221.

[8] H. Steiner, M. Jakusch, M. Kraft, M. Karlowatz, B. Mizaikoff, T. Baumann, R. Niessner, W. Konz, A. Brandenburg, K. Michel, C. Boussard-Pledel, B. Bureau, J. Lucas, Y. Reichlin, A. Katzir, N. Fleischmann, K. Staubmann, R. Allabashi, J. M. Bayona: In Situ Sensing of Volatile Organic Compounds in Groundwater: First Field Tests of a Mid-Infrared Fiber-Optic Sensing System, Society for Applied Spectroscopy, 57, 2003, 124-130.

[9] B. Obrovski, J. Bajić, I. Mihajlović, M. Vojinović Miloradov, B. Batinić, M. Živanov: Colorimetric fiber optic probe for measurement of chemical parameters in surface water, Sensors and Actuators B: Chemical, 228, 2016, 168-173.